## System Overview
System | Signal measured | Sensing hardware | Acquisition interface | Computer connection | Typical outputs | Unity control | Important note | Link
Wii Balance Board | Vertical loads; calculated CoP AP and ML | Four corner load sensors | WBB internal electronics | Bluetooth via compatible software | 4 corner loads, CoP X/Y, timestamps | Body lean moves car / player | CoP is computed from corner loads; log timestamps | https://github.com/tomvredeveld/wii-balance-board-recorder
BioRadio sEMG | Muscle electrical activity (voltage) | Surface electrodes on target muscle | BioRadio EMG-compatible inputs | BioRadio wireless link with bluetooth / supported acquisition software | Raw EMG, filtered EMG, RMS/MAV, activation | Muscle activation controls speed, direction, actions | Confirm model, channels, sample rate, SDK/software availability | https://www.glneurotech.com/products/bioradio/
FSR + Arduino | Pinch force proxy (ADC / calibrated N) | FSR e.g. FlexiForce; interface resistor/circuit | Arduino ADC | USB serial | ADC count, voltage, filtered force proxy | Pinch raises/lowers ball |  | https://www.tekscan.com/blog/flexiforce/differences-between-force-sensitive-resistor-technologies-assessing-linearity-drift

## Hardware Setup
System | Hardware required | Placement / wiring | Setup steps | Safety / checks | Sampling guidance | What to record
Wii Balance Board | WBB, batteries, Bluetooth-capable PC | Both feet on board; stable level floor | Pair Bluetooth; verify four corner loads; zero with no load; capture quiet standing | Clear surroundings; spotter for standing tasks | Measure actual rate from timestamps; avoid assuming fixed rate | Raw corner loads, calculated CoP, timestamps, trial labels
BioRadio sEMG | BioRadio, electrodes, leads, PC, acquisition software | Electrodes over muscle belly according to placement guidance; reference as specified by device | Check device/software; clean skin; connect channels; verify rest/contraction trace | Use approved isolated physiological acquisition equipment; follow device instructions | Choose EMG-appropriate sampling rate based on device capability; record actual value | Raw voltage, timestamps, muscle, side, posture, condition
FSR | Arduino interface with FSR, Arduino, suitable interface circuit, USB cable, PC | Sensor between finger/thumb pads with mechanical protection; appropriate divider/conditioning | Connect analog input; inspect ADC readings; record baseline and multiple force levels | Do not overload sensing area; confirm supply voltage and wiring | Choose and document a consistent acquisition rate | ADC count, voltage, calibration condition, timestamps

## For reference
System | Noise/artifacts to inspect | Time-domain analysis | Frequency-domain analysis | Online processing | Extracted features | Compare / evaluate
Wii Balance Board | Load jitter, spikes, drift, timing irregularities; natural sway is not noise | CoP trace, mean, SD, path, response | FFT/PSD for motion vs high-frequency components | Causal moving average and IIR low-pass; optional dead zone | CoP AP/ML, variability, filtered position | Raw vs filtered jitter, latency, steering accuracy
BioRadio sEMG | Motion artifacts, power-line interference, electrode contact noise | Waveform, rectified envelope, RMS, MAV | EMG spectrum / PSD; check 50 Hz interference | Appropriate band-pass, notch if needed, rectification and sliding RMS | RMS, MAV, onset/offset, activation duration | Rest vs contraction, false triggers, delay, task performance
FSR + Arduino | ADC quantization, electrical noise, mechanical hysteresis/drift | ADC/voltage traces, mean, SD, step response | FFT/PSD if useful for noise characterization | Median/outlier rejection and causal low-pass smoothing | Baseline, normalized pressure, calibrated force if available | Unfiltered vs filtered ball jitter, overshoot, delay

## Unity Integration
System | Suggested game | Input feature | Mapping | Calibration | Real-time data path | Failure handling | Game metrics
Wii Balance Board | Car racing / balloon dodge | Filtered CoP X,Y | X controls steering; Y optional second axis | Quiet stance center; reachable lean limits | Bluetooth acquisition -> Python bridge -> UDP/TCP -> Unity | Hold safe neutral on signal loss; clamp range | Collisions, target hits, time, movement delay
BioRadio sEMG | Car, balloon, football | Filtered EMG RMS / activation events | One muscle controls one action; two channels for two controls | Resting baseline and comfortable contraction reference | BioRadio software/SDK -> processing bridge -> Unity | Neutral state when connection drops; debounce triggers | False activations, response time, scores
FSR + Arduino | Pinch-controlled rising ball | Filtered normalized pinch reading | Increasing pinch moves ball up; releasing moves down | No-force baseline; comfortable pinch reference | Arduino USB serial -> PC processing / Unity | Clamp limits and avoid sudden jumps | Target error, stability, completion time